## Matirx_P_S_Total Control_German
11606 | Web-Rev.32 ('24.Aug.29) | [5 empty] | 152 | 155 | 154 | 154 | 155 | 155 | 155 | 155 | 155 | 29 | 23 | 23 | 17 | 25 | 25 | 27 | 21 | 155 | 11 | 11 | 55 | 86 | 33 | 32 | 155 | 155 | 12 | 18 | 20 | 167 | 167 | 11 | 10 | 166 | 38 | 28 | 28 | 27 | 166 | 166 | 168 | 168 | 168 | 167 | 167 | 18 | 168 | 20 | 166 | 78 | 82 | 145 | 99 | 99 | 166 | 102 | 166 | 166 | 166 | 70 | 74 | 168 | 168 | 168 | 168 | 168 | 168 | 168 | 27 | 63 | 63 | 166 | 168 | 145 | 171 | 168 | 168 | 168 | 168 | 168 | 168 | 67 | 65 | 55 | 57 | 57 | 59 | 59 | 155 | 56 | 166 | 166 | 16 | 166 | 19 | 166 | 28 | 28 | 31 | 31 | 168 | 9 | 26 | 168 | 168 | 32 | 166 | 19 | 36 | 36 | 30 | 53 | 46 | 46 | 10 | 33 | 32 | 22 | 22 | 8 | 21 | 18 | 31 | 28 | 24
Cell Count |  |  |  |  | Modified Information |  | System | System | System | System | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Volume | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Curved | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Endocavity | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Linear | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | Phased Array | CW | CW | CW | CW | CW | TEE | TEE | TEE
11606 | Classification | Name | Type | Disinfectanct Level | Manufacturer (Country) | Active Ingredient | Body Surface | Control Panel | Control Panel (HERA, RS), '23.Feb~ (V8~V5), '23.Jul~ | Monitor | CV1-8A | CV1-8AD | CV1-8AE | CMV1-10 | CMV1-10Z | 3DC2-6 | 3D2-6 | 3D2-6ET | V2-6 | 3D4-8 | 3D4-8ET | V4-8 | VE4-8 | VN4-8 | 3D4-9 | 3D4-9ES | V5-9 | EV3-10B | LV3-14A | V6-12 | EV2-10A | EV2-12 | C1-4EC | CS1-4 | SC1-6 | CA1-7A | CA1-7AD | C2-4/20 | C2-4ES | C2-5 | CA2-6BM | C2-6 | C2-6IC | CN2-8 | C2-8 | C2-8ED | CA2-8A | CA2-8AD | CF2-8 | CA2-9A | CA2-9AD | C3-7EP | CA3-10A | C4-9/10ED | CN4-9 | C5-8 | CF4-9 | CA4-10M | CA1-7S | CA1-7SD | EC4-9 | EC4-9IS | ER4-9 | EV4-9 | EVN4-9 | VR5-9 | EA2-11B | EA2-11AR | EA2-11AV | EA2-11ARD | EA2-11ARE | EA2-11AVD | EA2-11AVE | miniER7 | E3-12A | LA2-9A | L4-7 | L3-8 | L3-12A | L3-22 | LA2-9S | LA2-16S | LA3-14A | LA3-14AD | LA3-16A | LA3-16AD | LF5-13 | LA3-16AI | LS6-15 | LA3-22AI | LA4-18A | LA4-18AD | LA4-18B | LA4-18BD | LM2-18 | LM4-15B | LN5-12/40 | LN5-12 | LF5-12 | L5-12/50 | LS5-13 | L5-13/50 | L5-13 | L5-13IS | L7-16 | L7-16IS | LA2-14A | LA2-9R Vet | P1-4 | PA1-5A | PA1-5AE | PM1-6A | PN2-4 | P2-4AH | P2-4BA | PE2-4 | SP3-8 | PA3-8B | P3-8 | P3-8CA | PA3-9B | PA4-12B | P4-12 | DP2B | CW2.0 | CW 4.0 | CW6.0 | DP8B | TA2-9 | MMPT3-7 | MPT3-7
56 | Disinfectant | 3M SURFACE DISINFECTANT CLEANER WIPES | Wipes | Low | 3M (Canada) | Quaternary Ammonium(N-Alkyl), Alcohol | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
4 | Disinfectant | Accel INTERVention Wipes | Wipes | High | Diversey (Canada) | Hydrogen peroxide | ○ | X | X | ○
4 | Disinfectant | Accel Prevention Wipes | Wipes | High | Diversey (Canada) | Hydrogen peroxide | X | ○ | ○ | ○
56 | Disinfectant | Acrylan | Liquid | Low | Antiseptica (Germany) | Quaternary Ammonium(N-Alkyl) |  |  |  |  | ○ | ○ | ○ | ○ | ● | ○ | [7 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○
4 | Disinfectant | ALI-FLEX RTU | Liquid | Low | Lalema (Canada) | Sodium Hypochlorite | X | X | ○ | ○
2 | Cleaner | Alkazyme | Liquid | Cleaner | Alkapharm (France) | Quaternary Ammonium(N-Alkyl), Proteolytic Enzymes | [86 empty] | ○ | ○
36 | Disinfectant | Amity 3 step solution | Wipes | High | Amity (U.K) | Bis(3-aminoproply) dodecylamine, Didecyldimethyl Ammonium Chloride | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● |  |  |  | ● | ● |  |  |  | ● | [5 empty] | ● | ● | ● | ● | ● | [12 empty] | ● | ● |  | ● |  | ● | ● | [10 empty] | ● |  | ● | ● |  | ● |  | ● | [10 empty] | ●
56 | Cleaner | Amity 3 step solution: Detergent | Wipes | Cleaner | Amity (U.K) | Alkyl ethoxylate | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
56 | Disinfectant | Amity 3 step solution: HLD | Wipes | High | Amity (U.K) | Bis(3-aminoproply) dodecylamine, Didecyldimethyl Ammonium Chloride | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
36 | Cleaner | Amity 3 step solution: Rinse | Wipes | Cleaner | Amity (U.K) | Deionised water | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● |  |  |  | ● | ● |  |  |  | ● | [5 empty] | ● | ● | ● | ● | ● | [12 empty] | ● | ● |  | ● |  | ● | ● | [10 empty] | ● |  | ● | ● |  | ● |  | ● | [10 empty] | ●
56 | Disinfectant | Anios Excel Surface Premium | Liquid | Low | Anios (France) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ |  | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | X
4 | Disinfectant | ANIOSURF Premium NPC | Liquid | Low | Anios (France) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○
35 | Cleaner | Aniosyme DD1 | Liquid | Cleaner | Anios (France) | Quaternary Ammonium(N-Alkyl), Proteolytic Enzymes | [18 empty] | ○ | ○ |  | ○ | [7 empty] | ○ | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  | ○ | ○ | ○ | [7 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [19 empty] | ○ |  |  | ○ | ○
81 | Disinfectant | Anioxyde 1000 LD | Liquid | High | Anios (France) | Peracetic Acid | X | X | ○ | ○ | ● | ● | ● | ● | ● | ○ | [7 empty] | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  |  |  | ○ | ○ |  |  | X | ○ | X | X |  | X | X | ○ | ○ | ○ | ○ | ○ |  | ○ |  | X | ○ | ○ | ● | ○ | ○ | X | ○ | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X |  | ○ | ○ | X | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | ● | ● |  |  | ● | ○ | X | X |  | X |  | X |  |  | ○ | ○ | ○ |  |  | ○ | ○ |  | ○ | [7 empty] | ●
70 | Disinfection System | Antigermix S1 | UV | High | Germitec (France) | UV-C |  |  |  |  | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ |  |  | ★ |  |  |  | ★ | ★ |  |  |  | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ | ★ |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  | ★ | ★ | ★ |  | ★ |  | ★ | ★ |  | ★ | ★ | ★ |  | ★ | ★
57 | Disinfectant | Antigone wipes | Wipes | Low | Dan Mor Natural Products and Chemicals Ltd (Israel) | Alcohol, Quaternary Ammonium(N-Alkyl) | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
28 | Disinfection System | AquaCide® / Ethos Automated Ultrasound Probe Cleaner Disinfector | Liquid | High | CS Medical (USA) | peracetic acid |  |  |  |  | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ | [6 empty] | ★ | ★ | [14 empty] | ★ | ★ | ★ |  |  |  | ★ | [14 empty] | ★ | ★ | ★ | ★ | ★ | ★ | ★ | [7 empty] | ★ |  |  | ★ | ★ | ★ |  |  | ● |  |  |  |  | ★ | [11 empty] | ★ |  |  | ★ | ★
122 | Gel | Aquasonic 100 | Gel | Gel | Parker Laboratories, Inc. (USA) | NA | X | X | ○ | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ●
65 | Gel | Aquasonic Clear Ultrasound Gel | Gel | Gel | Parker Laboratories, Inc. (USA) | peracetic acid NA: NA; peracetic acid X: X; peracetic acid X: X; peracetic acid ○: ○; peracetic acid ○: ○; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ● |  |  | ● | ● |  | ● |  |  |  |  | ● | ● | ●
29 | Disinfectant | Asepti-Wipes II | Wipes | Low | Ecolab (U.K) | peracetic acid NA: Alcohol Quaternary Ammonium(N-Alkyl); ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ●: ● |  |  | ● | ●
3 | Disinfectant | Avert Sporicidal disinfectant cleaner | Liquid | Low | Diversey (Canada) | peracetic acid NA: Sodium hypochlorite; peracetic acid X: ○; peracetic acid ○: ○; peracetic acid ○: ○
4 | Cleaner | AZO detergent wipe | Wipes | Cleaner | Synergy Health Ltd (U.K) | peracetic acid NA: NA; peracetic acid X: X; peracetic acid X: ○; peracetic acid ○: ○; peracetic acid ○: ○
59 | Disinfectant | Azomax | Wipes | Low | Synergy Health Ltd (U.K) | peracetic acid NA: Quaternary Ammonium(N-Alkyl), Alcohol; peracetic acid X: X; peracetic acid X: ○; peracetic acid ○: ○; peracetic acid ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ○; ★ ○: ○; ★ ●: ●; ★ ○: ○; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ○; ★ ○: ○; ★ ●: ●; ★ ○: ○; ★ ●: ●; ★ ○: ○; ★ ○: ○; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ● |  |  | ● | ● |  | ○
53 | Disinfectant | Bacillol 30 Foam | Spray | Low | BODE Chemie GmbH (Germany) | peracetic acid NA: Alcohol Quaternary Ammonium(N-Alkyl); ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○ |  |  | ○ | ○ |  | ○
70 | Disinfectant | Bacillol 30 Sensitive Tissue | Wipes | Low | BODE Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ●
69 | Disinfectant | Bacillol 30 Tissue | Wipes | Low | BODE Chemie GmbH (Germany) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ●
4 | Cleaner | Bacillol AF | Wipes | Cleaner | BODE Chemie GmbH (Germany) | peracetic acid NA: Alcohol; peracetic acid X: X; peracetic acid X: X; peracetic acid ○: ○; peracetic acid ○: X
57 | Disinfectant | Biguacid S | Liquid | Low | Antiseptica (Germany) | peracetic acid NA: Quaternary Ammonium(N-Alkyl), Alcohol; peracetic acid X: X; peracetic acid X: X; peracetic acid ○: X; peracetic acid ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ○: ●; ★ ○: ●; ★ ○: X; ★ ○: X; ★ ○: X; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ○: X; ★ ○: X; ★ ○: X; ★ ○: X; ★ ○: X; ★ ○: ○; ★ ○: ○; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: X; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ○; ★ ○: ○; ★ ○: X; ★ ○: X; ★ ○: X; ★ ○: X; ★ ●: ● |  |  | ● | ● |  | X
4 | Cleaner | BIOSPOT* | Tablet | Cleaner | HYDRACHEM (U.K) | peracetic acid NA: Sodium Dichloroisocyanurate; peracetic acid X: ○; peracetic acid X: ○; peracetic acid ○: ○; peracetic acid ○: ○
2 | Disinfectant | Bleach 5.25% | Liquid | Low | Worldwide Production (Nation_wide) | peracetic acid NA: Sodium Hypochlorite; ★ ○: ○ | ○
61 | Disinfectant | CaviCide | Spray | Low | Metrex (USA) | peracetic acid NA: Alcohol Quaternary Ammonium(N-Alkyl); peracetic acid X: X; peracetic acid X: X; peracetic acid ○: ○; peracetic acid ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ●: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ●: ○ |  |  | ○ | ○ |  | ○
67 | Disinfectant | CaviWipes | Wipes | Low | Metrex (USA) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● |  |  |  | ● |  |  |  |  | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | X | X |  |  | X |  |  |  |  | X | X | ● | ● | ● | ● | ● |  | ● |  | X | ● | ● | ● | ● | ● | X |  | X | X | X | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● | ● | X | X |  | X |  | X |  |  |  |  | ● |  |  | ● | ● |  | ●
59 | Disinfectant | CaviWipes 1 | Wipes | Low | Metrex (USA) | peracetic acid NA: alcohol, Quaternary Ammonium(N-Alkyl); peracetic acid X: X; peracetic acid X: X; peracetic acid ○: X; peracetic acid ○: ○; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ●: ●; ★ ●: ●; ★ ●: ●; ● ○: ●; ★ ●: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ●: ● |  |  | ● | ● |  | ● |  |  |  |  | ● | ● | ●
5 | Disinfectant | CaviWipes AF | Wipes | Low | Metrex (USA) | peracetic acid NA: Quaternary Ammonium(N-Alkyl); ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●; ★ ○: ●
15 | Disinfectant | Chlorine 1000PPM | Liquid | High | Worldwide Production (Nation_wide) | peracetic acid NA: Chlorine compound; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○ | [5 empty] | ○
15 | Disinfectant | Chlorine Dioxide Solution | Liquid | High | vitalundfitmit100 Gmbh (Germany) | peracetic acid NA: Chlorine dioxide; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○; ★ ○: ○ | [5 empty] | ○
108 | Disinfectant | Cidex | Liquid | High | Advanced Sterilization Products (USA) | Glutaraldehyde | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○(50) | ○ |  |  | ○ | ○(50) | ○ | ○ | ○ | ○ |  | ○ | ○ | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | ○ | ● | ○(50) | ○(50) | ● | ● | ● | ● | ○ | ● | ● | ● | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ● | ● | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ○ | ○(50) | ● | ● | ○(50) | ● | ○ | ● |  |  |  |  | ● |  | ○ | ● | ● | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○(50) |  |  | ◆ | ◆ | ◆
117 | Disinfectant | Cidex OPA | Liquid | High | Advanced Sterilization Products (USA) | Ortho-phthalaldehyde | X | X | ○ | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(15) | ● | ○ | ○ | ○(15) | ◆ | ○(15) | ○(15) | ◆ | ● | ○ | ○ | ○ | ● | ● |  | ○ | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | ○ | ● | ○(15) | ○(15) | ● |  |  | ◆ | ◆ | ◆ | ◆ | ◆ | ○(15) | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● | ◆ | ○(15) | ○(15) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ● | ○(15) | ● | ● | ○(15) | ● | ○ | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○(15) | ○(15) | ○ | ○(15) | ○ | ◆ | ◆ | ◆
111 | Disinfectant | Cidex Plus | Liquid | High | Advanced Sterilization Products (USA) | Glutaraldehyde | X | X | ○ | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ● | ○ | ○ | ○(25) | ○(25) |  |  | ● | ● | ○ | ○ | ○ | ● | ● | ○ | ○ | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● |  | ● | ○(25) | ○(25) | ● | ○ | ○ | ● | ○ | ● | ● | ● | ○(25) | ○(25) | ● | ● | ● | ● | ● | ● | ● | ○ | ○(25) | ○(25) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ● | ○(25) | ● | ● | ○(25) | ● | ○ | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ | ○(25) | ○(25) | ○ | ○(25) | ○
79 | Cleaner | Cidezyme | Liquid | Cleaner | Advanced Sterilization Products (USA) | Proteolytic Enzymes |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ● |  |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | X | ● | ● | ● | ● | ○ | ○ | ● | ○ | ● | ● | ● | ○ | ○ | ◆ | ◆ | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  | ○ | ● | ● |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆
78 | Disinfectant | Cleanisept Wipe forte | Wipes | Low | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ |  |  | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●
78 | Disinfectant | Cleanisept Wipe forte Maxi | Wipes | Low | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | ○ | X | X | ○ | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ |  |  | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●
83 | Disinfectant | Cleanisept Wipe maxi | Wipes | Low | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | X | ○ | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ●(15) | ● |  |  |  | ● | ●(15) | ●(15) | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ◆ to ●
83 | Disinfectant | Cleanisept Wipes | Wipes | Low | Dr.Schumacher GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ●(15) | ● |  |  |  | ● | ●(15) | ●(15) | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ◆ to ●
1 | Gel | Clear Image | Gel | Gel | NEXT Medical Products (USA) | NA | [92 empty] | ○
56 | Disinfectant | Clinell Alcohol Wipes | Wipes | Low | GAMA Healthcare Ltd (U.K) | Isopropyl Alcohol 70% | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  |  |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
0 | Disinfectant | Clinell Clorox Wipes | Wipes | Low | GAMA Healthcare Ltd (U.K) | Sodium Hypochlorite
55 | Cleaner | Clinell Detergent Wipes(Yellow) | Wipes | Cleaner | GAMA Healthcare Ltd (U.K) | Lauryl Polyglucose, Butoxydiglycol | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
58 | Disinfectant | Clinell Peracetic Acid Wipes | Wipes | High | GAMA Healthcare Ltd (U.K) | Citric acid | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  |  |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ● to ●; ● to ●; ◆ to ●
117 | Disinfectant | Clinell Universal Santising wipes | Wipes | Low | GAMA Healthcare Ltd (U.K) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [23 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | bar: ● to ●
19 | Disinfectant | Clinell Universal Spray | Spray | Low | GAMA Healthcare Ltd (U.K) | Quaternary Ammonium(N-Alkyl) | [20 empty] | ○ | ○ | [11 empty] | ○ |  |  |  |  | ○ | ○ | [31 empty] | ○ | [18 empty] | ○ | ○ |  | ○ |  | ○ | [10 empty] | ○ | bar: ●; ● to ○
117 | Disinfectant | Clinell Universal wipes | Wipes | Low | GAMA Healthcare Ltd (U.K) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [23 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | bar: ● to ●
68 | Disinfectant | Clorox Healthcare® Bleach Germicidal Wipes | Wipes | Low | Clorox (USA) | Sodium hypochlorite | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ◆ to ●
71 | Disinfectant | Clorox Healthcare® Hydrogen Peroxide Cleaner Disinfectant Wipes | Wipes | Low | Clorox (USA) | Hydrogen Peroxide | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  |  | ● | ● | ○ | ○ |  |  |  | ○ | ○ |  |  | X |  |  |  |  | X | X | ● | ● | ● | ○ | ○ |  | ● |  | X | ○ | ○ | ● | ● | ● | X |  | X | X | X |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ | X | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ● | ● |  |  |  | ● | ● | ○ | ○ | X | X |  | X |  | X |  |  |  |  | ● |  |  | ● | ● | ● | X
4 | Disinfectant | Clorox regular (Clorox regular 1 : water 10) | Liquid | Low | Clorox (USA) | Sodium Hypochlorite | ○ | ○ | ○ | ○
4 | Cleaner | CN'F | Liquid | Cleaner | Neo Medical (Korea) | Alcohol | X | ○ | ○ | ○
33 | Disinfection System | D45 | UV | High | UV Smart (Netherland) | UV-C | [33 empty] | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | [28 empty] | ★ |  |  | ★ | ★ |  |  | ★ | ★ | ★ | [9 empty] | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  |  | ★ | ★ |  | ★ | bar: ●; ●; ●; ● to ●; ◆ to ●
46 | Disinfectant | deconex 11 UNIVERSAL | Liquid | Low | Borer Chemie AG (Switzerland) | Sodium hypochlorite | ○ | X | ○ | ○ | [25 empty] | ○ | ○ |  |  | ★ |  |  |  |  | ★ | ★ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ★ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [9 empty] | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ○ |  |  | ○ | ○ |  | ★ | bar: ●; ●; ●; ●; ● to ●; ◆ to ●
52 | Disinfectant | Descogen-I | Liquid | High | Antiseptica (Germany) | Potassium peroxymonosulfate, Citric Acid |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | bar: ●; ●; ●; ● to ●; ◆ to ●
52 | Disinfectant | Descoton Extra | Liquid | High | Dr.Schumacher GmbH (Germany) | Glutaraldehyde |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | bar: ●; ●; ●; ● to ●; ◆ to ●
59 | Disinfectant | Detro San AF Wipes | Wipes | Low | Detrox (Turkiye) | Quaternary Ammonium(N-Alkyl) | X | X | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
66 | Disinfectant | Dispatch Hospital Cleaner Disinfectant Towels with Bleach | Wipes | Low | Clorox (USA) | Sodium hypochlorite | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ● to ●; ◆ to ●
4 | Disinfectant | Distel Disinfectant Wipes | Wipes | High | Tristel Solutions Ltd. (U.K) | Polymeric Biguanide Hydrochloride | X | X | ○ | ○
1 | Disinfectant | Durvet Chlorhexidine | Liquid | Cleaner | Durvet (USA) | Chlorhexidine gluconate | [101 empty] | X
57 | Disinfectant | Ecobreeze Wipe | Wipes | Low | World of Disinfectant (Russia) | Alcohol, Quaternary Ammonium(N-Alkyl) | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
4 | Disinfectant | Ecolab Actichlor Tablets* | Tablet | Low | Ecolab (U.K) | adipic acid, Sodium Carbonate | ○ | ○ | ○ | X
75 | Gel | EcoVue | Gel | Gel | HR® Pharmaceuticals, Inc. (USA) | NA | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [23 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | bar: ●; ●; ● to ●; ● to ●
69 | Disinfectant | ED Wipes | Wipes | Low | MH Healthcare (Korea) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● |  |  |  |  | ● | bar: ●; ● to ●; ○ to ●
79 | Cleaner | EMpower | Liquid | Cleaner | Metrex (USA) | Propylene glycol | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [25 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆ | bar: ●; ●; ●; ● to ●
63 | Cleaner | EmPower Foam | Liquid | Cleaner | Metrex (USA) | Polyethylene glycol | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [25 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | bar: ●; ●; ●; ● to ○; ◆ to ●
60 | Cleaner | Endozime | Liquid | Cleaner | Ruhof (USA) | Proteolytic Enzymes | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ● | ● | [23 empty] | ● |  |  | ○ | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ●; ● to ●
2 | Cleaner | Endozime Sponge | Sponge | Cleaner | Ruhof (USA) | Proteolytic enzymes | [122 empty] | ● |  | ◆
15 | Cleaner | Endozime Xtreme Power | Liquid | Cleaner | Ruhof (USA) | Proteolytic Enzymes | [33 empty] | ○ |  |  |  |  | ○ | ○ | [31 empty] | ○ | [18 empty] | ○ | ○ |  | ○ |  | ○ | [10 empty] | ○ | bar: ●; ● to ●
15 | Disinfectant | Envirocide | Spray | Low | Metrex (USA) | Alcohol Quaternary Ammonium(N-Alkyl) | [33 empty] | ○ |  |  |  |  | ○ | ○ | [31 empty] | ○ | [18 empty] | ○ | ○ |  | ○ |  | ○ | [10 empty] | ○ | bar: ●; ● to ●
34 | Disinfectant | EnviroSafe Wipes | Wipes | Low | 경창메딕스 (Korea) | Alkyl(ethylbenzyl)dimethylammonium chloride |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● | [9 empty] | ● | ● | ● | ● | ● | [27 empty] | ● |  |  | ● | ● | ● | ● | ● | ● | [7 empty] | ● | [11 empty] | ● |  |  | ● | ● | bar: ●; ◆ to ●
79 | Cleaner | Enzol® | Liquid | Cleaner | Advanced Sterilization Products (USA) | Proteolytic Enzymes |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ● |  |  |  | ◆ | ● |  |  |  | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ● | ● | ● | ● | ● |  | ○ |  | ○ | ● | ● | ○ | ◆ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● |  |  |  | ○ | ● | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ● |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ◆ |  | ○ | ◆ | ◆ |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ | ● | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆
49 | Disinfectant | Enzyme Pre-Cleaner Wipes | Wipes | Low | HP&C (Korea) | Sodium Benzoate |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ●; ● to ●; ◆ to ●
4 | Disinfectant | Enzysept PLUS Wipes | Wipes | Low | HP&C (Korea) | Benzalkonium Chloride | ○ | ○ | ○ | ○
4 | Disinfectant | Ethanol 80% | Liquid | Low | Worldwide Production (Nation_wide) | Ethanol 80% | X | ○ | ○ | X
33 | Disinfectant | gigasept | Liquid | High | Schülke & Mayr GmbH (Germany) | Succindialdehyde, formaldehyde | [10 empty] | X | X |  | X | X | [6 empty] | ○(30) | [7 empty] | X | X | [9 empty] | X | X | X | X | X | [27 empty] | X | X | X | X | X | X | X | X | X | [19 empty] | X |  |  | X | X | bar: ●; ●; ◆ to ●
76 | Disinfectant | gigasept FF | Liquid | High | Schülke & Mayr GmbH (Germany) | Succindialdehyde |  |  |  |  | X | X | X | X | X |  | X | X |  | X | X | ○ | ○(30) | X |  |  |  | ○(30) | ○(30) | ○(30) | X | X | ○ |  |  | ○ | ○ |  | ○ | ○ |  | ★ | ★ |  | ○ | ○ | X | X | X | X | X |  | X |  | ○ | ○(30) | ○(30) | X |  |  | ○ | ○ | ○ | ○ | ○ | ○(30) | ○(30) | X | X | X | X | X | X | X |  | ○(30) | ○(30) | ○ | X | X | X | X | X | X | X | X | X | [7 empty] | X | ○(30) | ○ | ○ |  | ○ |  | ○ |  |  |  |  | X |  |  | X | X |  | ○ | [10 empty] | ○(30) | ○(30) | ○ | ○(30)
46 | Disinfectant | gigasept pearls* | Powder | High | Schülke & Mayr GmbH (Germany) | Sodium percarbonate, Citric Acid | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ● | ● | [9 empty] | ● | ● | ● | ○ | ○ | [27 empty] | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | [7 empty] | ○ | [11 empty] | ● |  |  | ● | ● | bar: ●; ● to ●; ○ to ●
66 | Disinfectant | gigasept® AF | Liquid | Low | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) |  |  |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ |  |  |  |  | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ○(30) | ○(30) | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(30) | ○(30) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | ○(30) | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | [10 empty] | ○(30) | ○(30) |  | ○(30)
52 | Disinfectant | gigasept® FF (new) | Liquid | High | Schülke & Mayr GmbH (Germany) | Succindialdehyde |  |  |  |  | X | X | X | X | X | [8 empty] | X | [6 empty] | X | X |  |  |  | X | X |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | X |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | bar: ●; ●; ●; ● to ●; ◆ to ●
53 | Disinfectant | gigasept® PAA concentrate | Liquid | High | Schülke & Mayr GmbH (Germany) | Hydrogen peroxide Acetic Acid | X | X | ○ | ○ | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ |  |  |  | ★ |  |  | ◆/★ | ◆/★ | [7 empty] | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ |  |  |  | ★ |  | ★ |  |  | ★ |  |  | ★ |  | ★ | ★ | ★ |  |  | ◆/★ | ◆/★ | ◆/★ | ◆/★ | ◆/★ | ◆/★ | ◆/★ | ★ |  |  | ★ |  | ★ | ★ | ★ |  |  | ★ | ★ | ★ | [7 empty] | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  |  | ★ | ★ |  | ★ | [15 empty] | ◆ | ◆ | ◆
57 | Disinfectant | gigasept® PowerSET3 | Wipes | High | Schülke & Mayr GmbH (Germany) | Hydrogen Peroxide / Acetic Acid / Peracetic Acid | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
55 | Disinfectant | HEXANIOS G+R | Liquid | Low | Anios (France) | Quaternary Ammonium(N-Alkyl) | ○ | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | X | bar: ●; ●; ●; ●; ● to ●; ◆ to ●
10 | Disinfectant | Hibitane | Liquid | High | Schering-Plough Animal Health Ltd. (New Zealand) | Chlorhexidine gluconate | [46 empty] | ●: ○(30); ●: ○(30); ○: ○(30); ○: ○(30) |  | ○(30) | ○(30) | [18 empty] | ○(30) | [27 empty] | ○(30) | ○(30) |  | ○(30)
58 | Disinfectant | HIS Wipes | Wipes | Low | Jun Medicare (Korea) | Choroalkylbenzyldimethylammonium | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  |  |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ● to ●; ● to ●; ◆ to ●
57 | Disinfectant | HMC NF* | Powder | High | MDD COMPANY GmbH (Switzerland) | Sodium percarbonate |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  | ○ | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | bar: ●; ●; ●; ● to ●
57 | Disinfectant | HORT Active Sporicidal Wipes | Wipes | Low | Polysept (Russia) | Hydrogen Peroxide | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
2 | Cleaner | Hydrichloric acid fuming 37% | Liquid | Cleaner | Millipore / Merck (다국적기업) | Hydrochloric acid | [122 empty] | ● | ●
68 | Disinfection System | Hypernova Chronos | UV | High | Germitec (France) | UV-C |  |  |  |  | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ |  |  | ★ |  |  |  | ★ | ★ |  |  |  | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | [23 empty] | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ | ★ |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  | ★ | ★ | ★ |  | ★ |  | ★ | ★ |  | ★ | ★ | ★ |  | ★ | ★ | bar: ●; ● to ●; ●; ● to ○; ◆ to ●
4 | Disinfectant | Incidin Foam | Spray | Low | Ecolab (U.K) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○
4 | Cleaner | Incidin Liquid | Liquid | Cleaner | Ecolab (U.K) | Alcohol | X | X | ○ | ○
63 | Disinfectant | Incidin oxyfoam | Foam | Low | Ecolab (U.K) | Hydrogen Peroxide | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ◆ to ●
49 | Disinfectant | Incidin oxyfoam S | Foam | Low | Ecolab (U.K) | Hydrogen Peroxide | X | ○ | ○ | ○ | [25 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ● to ●; ◆ to ●
68 | Disinfectant | Incidin oxywipe | Wipes | Low | Ecolab (U.K) | Hydrogen Peroxide | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  |  | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [23 empty] | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ● | bar: ●; ● to ●; ◆ to ●
64 | Disinfectant | Incidin oxywipe S | Wipes | Low | Ecolab (U.K) | Hydrogen Peroxide | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ◆ to ●
0 | Disinfectant | Incidin OxyWipes NG | Wipes | Low | Ecolab (U.K) | Hydrogen peroxide
60 | Disinfectant | ISO Wipe | Wipes | Low | Halyard (Australia) | Isopropyl alcohol 70% | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ○ to ●; ● to ●; ◆ to ●
96 | Cleaner | Isopropyl alcohol 70% | Liquid | Cleaner | Worldwide Production (Nation_wide) | Isopropyl alcohol 70% | X | X | ○ | ○ | ● | ● | ● | ● | ● | ○ | [7 empty] | ● |  |  |  | ○ |  |  | ● | ● |  |  | ○ | ● | ● |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | X | ● | ○(20) | ○(20) | ● | ● | ● | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ● |  | ○(15) | ○(15) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● |  | ● |  | ● | ○(15) | ○(15) | ○ | ○ | ● |  | ○ | ● | ● | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(15) | ○(15) |  | ○(15) |  | ● | ●
78 | Cleaner | Klenzyme | Liquid | Cleaner | Steris (USA) | Proteolytic Enzymes |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  | ○ |  |  | ○(10) | ○(10) |  |  | ○ | ○ |  |  |  | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | X | ● |  |  | ● |  |  | ● | ○ | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ |  |  |  | ○ | ○ | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  | ○ | ● | ● |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ | ● | ○ | ○ |  |  |  |  | ○
54 | Disinfection System | Lumicare ONE | UV | High | LUMICARE (Australia) | UV-C |  |  |  |  | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ | [6 empty] | ★ | ★ |  |  |  | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | [26 empty] | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  |  | ★ |  |  |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  |  | ★ | ★ |  | ★ | bar: ●; ●; ● to ●; ● to ●; ◆ to ●
4 | Disinfectant | Lysoform® | Liquid | High | Lysoform (Germany) | Formaldehyde | X | X | ○ | ○
0 | Cleaner | Matrix Wipes | Wipes | Cleaner | Whiteley Corporation (Australia) | Unique Biofilm Removing Detergent
88 | Cleaner | McKesson Alcohol Prep Pad Sterile | Wipes | Cleaner | McKesson Medical-Surgical Inc. (USA) | Alcohol |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ● | ● |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | X | ● | ○(20) | ○(20) | ● | ● | ● | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ● |  | ○(15) | ○(15) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ● | ● |  | ● |  | ● | ○(15) | ○(15) | ○ | ○ | ● |  | ○ | ● | ● | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(15) | ○(15) |  | ○(15)
43 | Disinfectant | MEDI CLEAN | Liquid | Low | UBM Cleaning Solution INC (Korea) | Quaternary Ammonium(N-Alkyl), Alcohol | [29 empty] | X | X |  |  | ○ |  |  |  |  | ○ | ○ | X | X | X | X | X | [26 empty] | ○ | X | X | X | X | X | X | X | X | X | [9 empty] | ○ | ○ |  | ○ |  | ○ |  |  |  |  | X |  |  | X | X |  | ○ | bar: ●; ●; ●; ● to ●; ◆ to ●
53 | Disinfectant | Medilox | Liquid | High | Soosan CMC (Korea) | Hypochlorous Acid | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● | [7 empty] | X |  |  |  |  | X | X | ○ | ○ | ○ |  |  |  | ○ |  | X |  |  | ○ | ● | ● | X |  | X | X | X |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | X |  | ○ | X | ○ |  |  | ○ | ○ | ○ |  |  | ● | ● | ● |  |  | ● |  | X | X |  | X |  | X |  |  |  |  | ○ |  |  | ○ | ○ |  | X
52 | Disinfectant | Medipal Alcohol Disinfectant Wipes | Wipes | Low | Pal International limited (U.K) | Alcohol |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | bar: ●; ●; ●; ● to ●; ◆ to ●
56 | Disinfectant | Medipal Chlorine Disinfectant Wipes | Wipes | Low | Pal International limited (U.K) | Sodium chloride | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  |  |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
54 | Disinfectant | Medipal Disinfectant Wipes | Wipes | Low | Pal International limited (U.K) | Quaternary Ammonium(N-Alkyl), Alcohol | X | X | ○ | ○ | [25 empty] | ● | ● |  |  | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | [23 empty] | ○ |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ● |  | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ● to ●; ◆ to ●
57 | Disinfectant | Mediwipes DM | Wipes | Low | Medilab Sp. Z.o.o (Poland) | Alcohol Quaternary Ammonium(N-Alkyl) | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ● to ●; ● to ●; ◆ to ●
4 | Cleaner | Meliseptol rapid | Liquid | Cleaner | B. Braun Medical AG (Switzerland) | Alcohol | X | X | ○ | ○
62 | Disinfectant | Meliseptol Wipes Sensitive | Wipes | Low | B. Braun Medical AG (Switzerland) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ◆ to ●
4 | Disinfectant | Melsept SF | Liquid | Low | B. Braun Medical AG (Switzerland) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○
89 | Disinfectant | Metricide | Liquid | High | Metrex (USA) | Glutaraldehyde |  |  |  |  | ○ | ○ | ○ | ○ | ○ |  | X | X |  | ○ | ○ | ○ | ○(50) | ○ |  |  | ○ | ○(50) | ○(50) | ○(50) | ○ | ○ |  | ○ |  | ○ | ○ |  |  | X |  | ○ | ○ |  | X | X | ● | ● | ● | ○ | ○ |  | ● |  | X | ○(50) | ○(50) | ● |  |  | X |  | X | X | X | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ● |  | ○(50) | ○(50) | X | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | [5 empty] | ○ | ○ | ○ | ○(50) | X | X | ○(50) | X | ○ | ○ |  |  |  |  | ● |  | ○ | ● | ● |  | ○ |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ |  | ○(50) |  | ◆ | ◆ | ◆
45 | Disinfectant | Metricide 28 | Liquid | High | Metrex (USA) | Glutaraldehyde | [15 empty] | ○ | [5 empty] | ○ | [7 empty] | ○ | ○ |  |  |  | ○ | [5 empty] | ● | ● | ● | ○ | ○ | [27 empty] | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | [11 empty] | ○ | [5 empty] | ● |  |  | ● | ● |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ◆ | ◆ | ◆ | bar: ●; ●; ○ to ●
68 | Disinfectant | Metricide 30 | Liquid | High | Metrex (USA) | Glutaraldehyde |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ | [26 empty] | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● |  |  | ○ |  | ○ | ○ |  | ○ |  | ○ | ○ | ○ |  |  | ● |  |  | ● | ● |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | bar: ●; ●; ●; ● to ●; ◆ to ●
78 | Disinfectant | Metricide OPA Plus | Liquid | High | Metrex (USA) | Ortho-phthalaldehyde | X | X | ○ | ○ | ● | ● | ● | ● | ● | ○ | [7 empty] | ● |  |  | ○ | ○ |  |  | ◆ | ● |  |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● | ◆ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● |  |  |  | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | X | X | ● | ● | ● |  |  | ● |  | ○ | ○ |  | ○ |  | ○ | X | X | X | X | ◆ |  |  | ● | ● |  | ○ | [15 empty] | ◆ | ◆ | ◆
71 | Disinfectant | Metricide Plus 30 | Liquid | High | Metrex (USA) | Glutaraldehyde | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [5 empty] | ● | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | [15 empty] | ◆ | ◆ | bar: ●; ● to ●
24 | Disinfectant | MetriGuard | Spray | Low | Metrex (USA) | Alcohol Quaternary Ammonium(N-Alkyl) |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ | [7 empty] | ○ |  |  |  |  | ○ | ○ | [31 empty] | ○ | [16 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ | [10 empty] | ○ | bar: ●; ● to ●
72 | Cleaner | Metrisponge | Sponge | Cleaner | Metrex (USA) | Propylene glycol |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [25 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  |  |  | ○ | bar: ●; ●; ●; ● to ●
81 | Cleaner | Metrizyme | Liquid | Cleaner | Metrex (USA) | Proteolytic Enzymes |  |  |  |  | ○ | ○ | ○ | ○ | ○ |  | X | X |  | X | X |  |  | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  | ○ | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | X | ● |  |  | ● |  |  | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ○ | ○ |  |  |  |  | ○ | ◆ | ◆ | ◆
59 | Disinfectant | Micro Kill Bleach Wipes | Wipes | Low | Medline (USA) | Sodium Hypochlorite | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  |  |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● |  |  |  |  | ● | ● | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
59 | Disinfectant | Mikrobac Tissues | Wipes | Low | BODE Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ●; ● to ●
66 | Disinfectant | Mikrozid AF Wipes | Wipes | Low | Schülke & Mayr GmbH (Germany) | Alcohol | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ●; ○ to ●; ● to ●; ◆ to ●
57 | Disinfectant | Mikrozid PAA Wipes | Wipes | Low | Schülke & Mayr GmbH (Germany) | Peracetic Acid / Hydrogen Peroxide / Acetic Acid | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
27 | Disinfectant | Mikrozid sensitive liquid | Liquid | Low | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ | [7 empty] | ○ |  |  |  |  | ○ | ○ | [31 empty] | ○ | [16 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ | [10 empty] | ○ | bar: ●; ● to ○
73 | Disinfectant | Mikrozid sensitive Wipes | Wipes | Low | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | bar: ●; ● to ●; ● to ●
28 | Disinfectant | Mikrozid universal liquid | Liquid | Low | Schülke & Mayr GmbH (Germany) | Alcohol | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● | [7 empty] | ○ |  |  |  |  | ○ | ○ | [31 empty] | ○ | [16 empty] | ● |  | ○ | ○ |  | ○ |  | ○ | [10 empty] | ○ | bar: ●; ● to ●
55 | Disinfectant | Mikrozid universal Wipes premium | Wipes | Low | Schülke & Mayr GmbH (Germany) | Alcohol | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | X |  |  |  |  | X | X | ● | ● | ● | ● | ● | [26 empty] | X | ● |  | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | X | X |  | X |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
76 | Disinfectant | Milton anti bacterial solution | Liquid | Low | Milton (France) | Sodium Hypochlorite | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ | ○ |  |  | ● | ● |  | ○ | ● |  | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ○ |  | ● | ● |  | ● |  | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ● |  | ● |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | bar: ● to ●; ●; ● to ●; ◆ to ●
57 | Disinfectant | Mirodez Wipes | Wipes | Low | World of Disinfectant (Russia) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
45 | Gel | Natural Image | Gel | Gel | Sonotech Inc. (USA) | NA |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ○ | ○ |  |  |  | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | [26 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ● | ● | ● |  |  | ● | [11 empty] | ○ |  |  | ○ | ○ |  |  | ○ | [6 empty] | ● | [6 empty] | ○ | bar: ●; ●; ◆ to ●
58 | Disinfectant | NDP SurfaPlus | Spray | Low | Vesismin Health (Spain) | Isopropyl alcohol | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  |  | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
3 | Cleaner | Neodisher® MediClean forte | Wipes | Cleaner | Dr.Weigert (Germany) | Trisodium nitrilotriacetate | [122 empty] | ◆ | ◆ | ◆
1 | Disinfectant | Nolvasan®Solution Chlorhexidine diacetate | Liquid | Cleaner | Zoetis (USA) | Chlorhexidine diacetate | [101 empty] | ○
55 | Disinfectant | Nuclean | Liquid | Low | Alkapharm (France) | Alcohol |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | X | [5 empty] | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ● to ●; ●; ●; ● to ●; ◆ to ●
56 | Disinfectant | Omnicide - FG2 | Liquid | High | Coventry Chemicals Limited (U.K) | Glutaraldehyde |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | X | X |  |  | ● |  |  |  | ○ | ● | ● | ● | ● | ● | X | X | [26 empty] | ● | X | ● | X | ● | X | X | ● | ● | ● | ○ | ○ | [5 empty] | ○ |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ● to ●; ●; ● to ●; ◆ to ●
36 | Disinfectant | Omnicide 14 | Liquid | High | Metrex (USA) | Glutaraldehyde | [29 empty] | ○ | ○ | [9 empty] | ● | ● | ● | ○ | ○ | [27 empty] | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | [19 empty] | ● |  | ○ | ● | ● |  |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | bar: ●; ●; ◆ to ●
39 | Disinfectant | Omnicide28 | Liquid | High | Metrex (USA) | Glutaraldehyde | [29 empty] | ○ | ○ |  |  |  | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | [27 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [19 empty] | ○ |  | ○ | ○ | ○ |  |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | bar: ●; ●; ◆ to ●
0 | Disinfectant | OPAL Disinfectant | Wipes | Low | Whiteley Corporation (Australia) | O-phthaldehyde
47 | Gel | Opti Lube Gel | Gel | Gel | Optimum Medical (U.K) | NA | [29 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ● to ●; ◆ to ●
60 | Disinfectant | OptiGerm | Spray | Low | Oleak (Brazil) | PHMB, Quarternary Ammonium | X | X | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ○ to ●; ● to ●; ◆ to ●
60 | Disinfectant | Oxivir 1 | Wipes | Low | Diversey (Canada) | Hydrogen Peroxide | X | X | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ● to ●; ● to ●; ◆ to ●
75 | Disinfectant | Oxivir TB Wipes | Wipes | Low | Diversey (Canada) | Hydrogen Peroxide | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ● | ● | [23 empty] | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● |  |  | ○ | ○ | ● | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ○ | ● | bar: ●; ● to ●; ○ to ●
4 | Disinfectant | Oxivir® Five 16 Concentrate | Liquid | High | Diversey (Canada) | Hydrogen Peroxide |  | X | ○ | ○ | [97 empty] | X
37 | Disinfectant | Peraflu D | Liquid | High | Velly Co.,Ltd (Korea) | Peracetic Acid |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ | [27 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | [11 empty] | ○ |  |  | ○ | ○ | bar: ●; ●; ◆ to ●
93 | Disinfectant | PeraSafe* | Powder | High | Dupont (Germany) | Citric Acid | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ○ | X | X |  | X | X | ○ | ○(15) | ● |  |  | ○(15) | ○(15) | ○(15) | ○(15) | ● | ● |  |  |  | ○ | ○ |  |  | X | ○ | ○ | ○ |  | X | X | ● | ● | ● | ○ | ○ |  | ● |  | X | ○(15) | ○(15) | ● | ○ | ○ | X | X | X | X | X | ○(15) | ○(15) | ● | ● | ● | ● | ● | ● | ● |  | ○(15) | ○(15) | X | X | ● | X | ● | X | X | ● | ● | ● | ○ | ○ | ● | ● | ● |  |  | ● | ○(15) | X | X |  | X | ○ | X |  |  | ○ | ○ | ● |  |  | ● | ● |  | X | [10 empty] | ○(15) | ○(15) |  | ○(15) |  | ◆ | ◆ | ◆
4 | Disinfectant | perform®* | Powder | High | Schülke & Mayr GmbH (Germany) | Pentapotassium | X | X | ○ | ○
3 | Cleaner | Prolystica® 2X Concentrate Enzymatic Presoak & cleaner | Wipes | Cleaner | Steris (USA) | Ethanolamine Ethoxylated alcohol Polyalkylene glycol Glycerine | [122 empty] | ◆ | ◆ | ◆
73 | Disinfectant | Protex disinfectant spray | Spray | Low | Parker Laboratories, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ |  |  | ○ |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [24 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ | ○ |  | ○ | ○ |  | ○ | bar: ●; ● to ●
63 | Disinfectant | Protex Pro Disinfection wipe | Wipes | Low | Parker Laboratories, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ● to ●
68 | Disinfectant | PROTEX ULTRA DISINFECTANT WIPES | Wipes | Low | Parker Laboratories, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ◆ to ●
4 | Disinfectant | Pursept AF | Liquid | Low | Schülke & Mayr GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | X | X
1 | Disinfectant | Quatricide® PV Second Generation | Liquid | Low | Pharmacal Research Laboratories (USA) | Quaternary Ammonium(N-Alkyl) | [101 empty] | X
48 | Disinfectant | Rapicide OPA/28 | Liquid | High | Medivators (USA) | Ortho-phthalaldehyde | [20 empty] | ○ | ○ | [7 empty] | X | X |  |  | ○ |  |  |  |  | ○ | ○ | X | X | X | X | X | [26 empty] | ○ | X | X | X | X | X | X | X | X | X | [8 empty] | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | X |  |  | X | X |  | ○ | bar: ●; ●; ●; ● to ●
3 | Disinfectant | Rapicide PA HLD | Liquid | High | Medivators (USA) | Peracetic acid | [122 empty] | ◆ | ◆ | ◆
6 | Disinfectant | Rely+On Virkon* | Powder | High | LANXESS Corporation (Australia) | Potassium peroxymonosulfate | [26 empty] | ○ | [41 empty] | ○ | [27 empty] | ○ | ○ | ○ | ○
72 | Disinfectant | Revital-Ox® RESERT XL HLD | Liquid | High | Steris (USA) | Hydrogen Peroxide |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [23 empty] | ○ | ○ | ○ | ○ | ○ | ● | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● |  |  | ● | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | [15 empty] | ◆ | ◆ | ◆ | bar: ●; ●; ● to ●
69 | Disinfectant | RHEOSEPT-WD plus Wipes | Wipes | Low | Wachendorff-Chemie GmbH (Germany) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ● | bar: ●; ● to ●; ● to ●; ◆ to ●
60 | Disinfectant | Rudocid | Wipes | Low | Medidez (Russia) | isopropyl alcohol, benzalkonium chloride | X | X | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ● | bar: ●; ●; ● to ●; ● to ●; ◆ to ●
48 | Disinfectant | Rudocid®ENDO | Wipes | Low | Medidez (Russia) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ○ | X | [25 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ● to ●; ● to ●; ◆ to ●
75 | Cleaner | saFecide Medical wipes | Wipes | Cleaner | Cosell Care (Korea) | Alcohol | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | bar: ●; ● to ●; ● to ●; ○ to ●
66 | Disinfectant | saFecide Sensitive wipes | Wipes | Low | Cosell Care (Korea) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●; ○ to ●
10 | Disinfectant | Salvanios pH10 | Liquid | Low | Anios (France) | Guanidinium acetate Quaternary Ammonium(N-Alkyl) | [86 empty] | ○ | ○ | [14 empty] | ○ | [5 empty] | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○
53 | Disinfectant | Sani Cloth active | Wipes | Low | Ecolab (U.K) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | [17 empty] | ● | [7 empty] | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [24 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [9 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ●; ● to ●; ○ to ●
105 | Cleaner | Sani-Cloth 70 | Wipes | Cleaner | PDI (U.K) | Alcohol | X | ○ | ○ | ○ | ● | ● | ● | ● | ● |  | ○ | ○ | ○ | ● | ● | ○ | ○ | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  | ○ |  | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [24 empty] | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ |  |  | ○ | ○ | ● | ○ | ○ | ○ |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ● | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ |  | ○ | bar: ●; ● to ●
73 | Disinfectant | Sani-Cloth Active | Wipes | Low | PDI (U.K) | Quaternary Ammonium(N-Alkyl) | ○ | X | ○ | ○ | ● | ● | ● | ● | ● |  |  |  | ● |  |  | ● |  | ● |  |  | ● | ● | ●(15) | ●(15) | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | [26 empty] | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ●; ● to ●
56 | Cleaner | Sani-Cloth Detergent Wipes | Wipes | Cleaner | PDI (U.K) | NA | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
101 | Disinfectant | Sani-Cloth HB | Wipes | Low | PDI (U.K) | Quaternary Ammonium(N-Alkyl) Tetrasodium EDTA | X | ○ | ○ | ○ | ● | ● | ● | ● | ● |  | X | X |  | X | X | ● | ●(30) | ● |  |  | ◆ | ● |  |  | ● | ● |  | ● |  | ● | ● |  |  | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | ● | ●(30) | ●(30) | ● | ◆ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ●(30) | ●(30) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ●(30) | ● | ● | ●(30) | ● | ● | ● |  |  |  |  | ◆ |  | ● | ◆ | ● |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ●(30) | ●(30) |  | ●(30) | ●
71 | Disinfectant | Sani-Cloth Plus | Wipes | Low | PDI (U.K) | Alcohol Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  | ● | ● | ● |  | ● |  | ● | ● |  | ● | ● | ● |  | ● | ●
55 | Disinfectant | Sani-Cloth Prime | Wipes | Low | PDI (U.K) | Isopropyl alcohol | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
3 | Disinfectant | Sani-Cloth Universal wipes | Wipes | Low | PDI (U.K) | NA: Quaternary Ammonium; X: X; ○: X; ○: ○
57 | Disinfectant | Sani-Cloth® AF Universal | Wipes | Low | PDI (U.K) | 2-Aminoethanol, D-D-A-C, Potassium, Phosphoric Acid | X | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  |  | ● |  |  |  |  | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● |  |  |  |  | ●
83 | Disinfectant | Sani-Cloth® AF3 Germicidal Disposable Wipe | Wipes | Low | PDI | Quaternary Ammonium(N-Alkyl) | X | X | X | X | ● | ● | ● | ● | ● | ● | [5 empty] | ● |  | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ●
87 | Disinfectant | Sani-Cloth® Bleach Germicidal Disposable Wipe | Wipes | Low | PDI (U.K) | Sodium Hypochlorite | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | [5 empty] | ● |  | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ●
109 | Gel | Scan | Gel | Gel | Parker Laboratories, Inc. (USA) | [13 empty] | ○ |  |  | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | bar: NA; ● to ●
56 | Disinfectant | SciCan Optim 33TB Surface Cleaner Disinfectant Solution | Liquid | Low | SciCan (Canada) | Hydrogen peroxide | X | X | X | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○
55 | Disinfectant | SciCan Optim 33TB Surface Cleaner Disinfectant Wipes | Wipes | Low | SciCan (Canada) | Hydrogen peroxide | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● |  |  | ○ |  | ○ | ○ | ○ |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● | ● |  | X
55 | Disinfectant | SCOP'ANIOS RC | Liquid | High | Anios (France) | Hydrogen peroxide, Peracetic Acid | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | X | X |  |  | X |  |  |  |  | X | X | ○ | ○ | ○ | ○ | ○ |  | ○ |  | X |  |  | ○ |  |  | X |  | X | X | X |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  | X | X |  | X |  | X |  |  |  |  | ○ |  |  | ○ | ○ |  | X
35 | Disinfectant | Sekusept Aktiv* | Powder | High | Ecolab (U.K) | Sodium percarbonate, Citric acid | X | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | X |  |  |  |  | X | X |  |  |  | ● | ● |  |  |  | X | [5 empty] | X |  | X | X | X | [12 empty] | X | ● |  | ● |  | ● | ● | [10 empty] | ○ |  | X | X |  | X |  | X | [10 empty] | X
4 | Disinfectant | Sekusept Extra | Liquid | High | Ecolab (U.K) | NA: Glutaraldehyde |  | X | X |  | X | X
2 | Cleaner | Sekusept Plus | Liquid | Cleaner | Ecolab (U.K) | NA: Glucoprotamin | [11 empty] | ○(30) | ○(30)
62 | Disinfectant | Septalkan | Wipes | Low | Alkapharm (France) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
34 | Disinfectant | Septiwipes | Wipes | Low | Dr.Schumacher GmbH (Germany) | NA: Quaternary Ammonium(N-Alkyl) | [6 empty] | ● | ●(15) |  |  |  |  | ● | ●(15) | ●(15) | [5 empty] | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  |  |  | ● | [9 empty] | ● | ● | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | [13 empty] | ● | [5 empty] | ● |  |  | ● | ●
2 | Disinfectant | SEPTOL LINGETTES | Wipes | Low | ACTEON (Italy) | NA: Quaternary Ammonium(N-Alkyl) | [77 empty] | X | X
4 | Cleaner | SideKick Disinfecting Wipes | Wipes | Cleaner | Stryker Medical (USA) | NA: Isopropyl alcohol; X: X; ○: ○; ○: ○; ○: ○
2 | Disinfectant | Soluscope PAA | Liquid | High | Ecolab (U.K) | NA: Hydrogen peroxide, Acetic Acid, Peracetic Acid | [113 empty] | ◆ | ◆
56 | Disinfectant | Sono Ultrasound Wipes | Wipes | Low | Advanced Ultrasound Solutions Inc. (USA) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● |  |  | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
35 | Gel | Sonogel | Gel | Gel | Sonogel Vertriebs GmbH (Germany) | [13 empty] | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ | ○ | ○ |  |  |  | ○ | [5 empty] | ● |  |  |  |  | ● | ● | [8 empty] | ● | ○ | ○ |  |  |  | ● | ● | ● | ● | ● | ○ | ○ | [8 empty] | ○ | ○ | ● |  |  | ● | [14 empty] | ○ | ● | ● | ○ | ● |  | ● | [10 empty] | ● | [10 empty] | ○ | ○ |  | ○ | bar: NA
1 | Disinfectant | Sporicidin® disinfectant solution | Liquid | Cleaner | Contec, Inc. (USA) | NA: Phenol | [92 empty] | ○
78 | Disinfectant | Sporox | Liquid | High | DSHealthcare Inc. (USA) | Hydrogen Peroxide | X | X | ○ | ○ | ● | ● | ● | ● | ● | X | X | X |  | X | X |  |  | ● | ○ | ○ |  |  |  |  | ● | ● |  |  | ○ | ○ | ○ | ○ |  | ○ | ○ |  |  |  | ○ | ○ | X | X | X | ○ | ○ | ○ | X | X | ○ |  |  | X |  |  | ○ |  | ○ | ○ | ○ |  |  | X | X | X | X | X | X | X | ○ |  |  | ○ | ○ | X | ○ | X | ○ | ○ | X | X | X | ○ | ○ | [5 empty] | ● |  | X | ○ |  | ○ |  | X | ○ | ○ | ○ | ○ | X |  |  | X | X | ○ | ○ | ○ |  |  | ○ | [10 empty] | ○
79 | Disinfectant | SporoxⅡ | Liquid | High | DSHealthcare Inc. (USA) | Hydrogen Peroxide | X | X | ○ | ○ | ● | ● | ● | ● | ● | X | X | X |  | X | X |  |  | ● | ○ | ○ |  |  |  |  | ● | ● |  |  | ○ | ● | ● | ○ |  | ● | ○ |  |  |  | ● | ● | X | X | X | ● | ● | ○ | X | X | ● |  |  | ○ |  |  | ● |  | ● | ● | ● |  |  | X | X | X | X | X | X | X | ○ |  |  | ● | ● | ○ | ● | X | ● | ● | X | X | X | ○ | ○ | [5 empty] | ● |  | ● | ● |  | ● |  | ● | ○ | ○ | ○ | ○ | X |  |  | X | X | ○ | ● | ○ |  |  | ○ |  |  |  | ● | [6 empty] | ○
45 | Disinfectant | Steranios 2% | Liquid | High | Anios (France) | NA: Glutaraldehyde | [11 empty] | ○ | ○ | [7 empty] | ○ | ○ | [9 empty] | ● | ● | ● | ○ | ○ |  | ● |  |  |  |  | ● | [7 empty] | ○ | ○ | ● | ● | ● | ● | ● | ● | ● |  |  |  |  | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | [5 empty] | ○ | ○ |  | ○ | [10 empty] | ● |  |  | ● | ● |  |  |  | ○ | ○ |  | ○ | ○ | ○ |  | ○ | ○ | [5 empty] | ◆ | ◆ | ◆
1 | Disinfectant | Steranios 2% ECS | Liquid | High | Anios (France) | NA: Glutaraldehyde | [115 empty] | ◆
1 | Disinfectant | Steranios 2% NG | Liquid | High | Anios (France) | NA: Glutaraldehyde | [115 empty] | ◆
15 | Disinfectant | Ster-Bac Blu | Liquid | Low | Ecolab (U.K) | NA: Quaternary Ammonium(N-Alkyl) |  | X | X |  | X | X | ○ | ○(50) | [10 empty] | ○ | [41 empty] | ○(50) | ○(50) | [18 empty] | ○(50) |  |  | ○(50) |  | ○ | [22 empty] | ○(50) | ○(50) |  | ○(50)
57 | Disinfectant | Sterisept Wipes | Wipes | Low | AS Chemi-Pharm (Estonia) | Quaternary Ammonium(N-Alkyl) | ○ | X | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  | ● | ● | ● | ● |  | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
4 | Disinfectant | Super Sani-Cloth Plus Wipes | Wipes | Low | PDI (U.K) | NA: Quaternary Ammonium(N-Alkyl); X: X; ○: ○; ○: ○; ○: ○
81 | Disinfectant | Super Sani-Cloth® Germicidal Disposable Wipe | Wipes | Low | PDI (U.K) | Alcohol Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● | ● |  |  | ● | ● |  |  | X |  |  |  |  | X | X | ● | ● | ● | ● | ● |  | ● |  | X | ● | ● | ● | ● | ● | X |  | X | X | X |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | X | X |  | X |  | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● |  |  |  |  | ● | ● | ●
24 | Disinfectant | SURFA' SAFE | Liquid | Low | Anios (France) | NA: Quaternary Ammonium(N-Alkyl); X: ○; ○: ○; ○: ○; ○: ○ | [11 empty] | ○ | ○ | [11 empty] | ○ |  |  |  |  | ○ | ○ | [8 empty] | ○ | [5 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [10 empty] | ○ | [17 empty] | ○ | ○ | ○ |  | ○ |  | ○ | [10 empty] | ○
1 | Disinfectant | Surfaces Hautes | Spray | High | Anios (France) | NA: Glutaraldehyde | [112 empty] | ○
28 | Disinfectant | Surfanios Premium | Liquid | Low | Anios (France) | NA: Quaternary Ammonium(N-Alkyl), Alcohol | [20 empty] | ○ | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ |  | ○ |  |  |  |  | ○ | [9 empty] | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [19 empty] | ○ |  |  | ○ | ○
61 | Disinfectant | SURFA'SAFE PREMIUM | Spray | Low | Anios (France) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ● | ● |  |  | ○ | ○ |  |  |  | ○ | ○ | ● | ● | ● | ● | ● |  | ● |  | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ |  |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | [7 empty] | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● | ● |  | ○
4 | Disinfectant | Tana Apesin spray | Spray | Low | tana Chemie GmbH (Germany) | NA: Quaternary Ammonium(N-Alkyl); X: X; ○: X; ○: ○; ○: ○
3 | Disinfection System | TD-5® (w/TD100) | Liquid | High | CS Medical (USA) | NA: Glutaraldehyde | [113 empty] | ◆ | ◆ | ◆
1 | Disinfection System | TD-8® (w/TD100) | Liquid | High | CS Medical (USA) | NA: Ortho-phthalaldehyde | [113 empty] | ●
2 | Cleaner | TEEZyme Sponge | Sponge | Cleaner | CS Medical (USA) | NA: Proteolytic enzymes | [113 empty] | ● |  | ◆
4 | Cleaner | terralin liquid | Liquid | Cleaner | Schülke & Mayr GmbH (Germany) | NA: Alcohol; X: X; ○: ○; ○: ○; ○: ○
31 | Disinfectant | terralin PAA | Liquid | High | Schülke & Mayr GmbH (Germany) | NA: Hydrogen peroxide, Acetic Acid, Peracetic Acid; X: X; ○: X; ○: X | [20 empty] | X | X | [9 empty] | X | X | X | X | X |  | X |  |  |  |  | X | [9 empty] | X | X | X | X | X | X | X |  |  |  |  | X | X | X | X | X | X | X | X | X | [19 empty] | X |  |  | X | X
95 | Cleaner | Transeptic® Cleaning Solution | Spray | Cleaner | Parker Laboratories, Inc. (USA) | Isopropyl alcohol 70% | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [7 empty] | ○ |  |  | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ |  |  | ○ |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ | ○(20) | ○(20) | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(15) | ○(15) | ○ | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ○ | ○ | ○ | ○(15) | ○ | ○ |  | ○ |  | ○ | ○(15) | ○(15) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(15) | ○(15) |  | ○(15)
85 | Disinfectant | Tristel Duo ULT | Foam | High | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | X | ○ | X | ○ | ● | ● | ● | ● | ● | ○ | [7 empty] | ● | ○ | ○ | ◆ | ◆ |  |  | ◆ | ● |  |  |  | ○ | ○ |  |  | ● | ○ | ○ | ○ | X | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● |  | ● | ● | ○ | ○ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● | ◆ |  |  | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ○ | X | X |  | X |  | X | ○ | ○ | ○ | ○ | ● |  |  | ● | ● |  | X |  |  |  | ○ | ●
66 | Disinfectant | Tristel Fuse for Stella | Liquid | High | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | ○ | ○ | ○ | ○ | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ |  |  |  | ★ |  |  | ★ | ★ |  |  |  | ● | ● | ● | ● | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | ● | ● |  | ★ |  | ★ |  |  | ★ | ● | ● | ★ |  | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ● | ● | ★ | ● | ★ | ★ | ★ | ● | ● | ★ | ★ | ★ | ● | ● | ● |  |  |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  |  | ★ | ★ |  | ★
77 | Cleaner | Tristel Pre-clean wipe | Wipes | Cleaner | Tristel Solutions Ltd. (U.K) | Alcohol | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● |  |  |  | ● | ● |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ | ● | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● | ● |  | ● |  |  |  |  | ● | [10 empty] | ◆ | ◆
72 | Cleaner | Tristel Rinse Wipes | Wipes | Cleaner | Tristel Solutions Ltd. (U.K) | Deionised water | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  |  | ● | ● | ● | ● | ● | ○ | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | [15 empty] | ◆ | ◆
69 | Disinfectant | Tristel Solo Wipes | Wipes | Low | Tristel Solutions Ltd. (U.K) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ●
84 | Disinfectant | Tristel Sporicidal Wipe | Wipes | High | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | ○ | [7 empty] | ● | ○ | ○ |  | ○ |  |  | ● | ● |  |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  | ● | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ |  | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● |  | ○ | ○ |  | ○ |  | ○ | ○ | ○ | ○ | ○ | ● |  |  | ● | ● |  | ○ |  |  |  | ○ | ● | [10 empty] | ◆ | ◆
63 | Disinfectant | Tristel Trio Wipes System | Wipes | High | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | ○ | X | X | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ |  |  |  | ○ |  |  | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ● | ● | ● | ○ | ○ |  | ● |  | ○ |  |  | ● |  |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ● | ● | ● | ● | ● | ● | ● |  |  |  | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | [5 empty] | ● | ● | ○ |  | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ● |  |  | ● | ● |  | ○ | [15 empty] | ◆ | ◆ | ◆
85 | Disinfectant | Tristel ULT | Foam | High | Tristel Solutions Ltd. (U.K) | Chlorine dioxide | X | ○ | X | ○ | ● | ● | ● | ● | ● | ○ | [7 empty] | ● | ○ | ○ | ◆ | ◆ |  |  | ◆ | ● |  |  |  | ○ | ○ |  |  | ● | ○ | ○ | ○ | X | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● |  | ● | ● | ○ | ○ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ● | ◆ |  |  | ○ | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ○ | X | X |  | X |  | X | ○ | ○ | ○ | ○ | ● |  |  | ● | ● |  | X |  |  |  | ○ | ●
80 | Cleaner | trophon® companion cleaning wipes | Wipes | Cleaner | Nanosonics Limited (Australia) | Quaternary Ammonium(N-Alkyl) | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  |  | ● | ● |  | ● | ● |  | ● |  | ○ | ○ |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ● | ○ | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● | ● | ● |  |  |  | ● | ● | ● | ●
95 | Disinfection System | trophon® EPR | Vapor | High | Nanosonics Limited (Australia) | Hydrogen Peroxide |  |  |  |  | ★ | ★ | ★ | ★ | ★ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ★ | ◆ | ◆ | ◆ | ◆ | ○ | ○ | ◆ | ★ |  | ○ | ○ | ○ | ○ |  |  | ● |  | ○ | ○ |  | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● | ○ | ○ | ● | ◆ | ● | ◆ |  | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ★ | ◆ | ○ | ○ | ● | ○ | ● | ● | ★ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ★ |  | ● | ● |  | ● |  | ● |  |  | ○ | ○ | ◆ |  |  | ● | ★ | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○
95 | Disinfection System | trophon®2 | Vapor | High | Nanosonics Limited (Australia) | Hydrogen Peroxide |  |  |  |  | ★ | ★ | ★ | ★ | ★ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ★ | ◆ | ◆ | ◆ | ◆ | ○ | ○ | ◆ | ★ |  | ○ | ○ | ○ | ○ |  |  | ● |  | ○ | ○ |  | ● | ● | ● | ● | ● | ○ | ○ |  | ● |  | ● | ○ | ○ | ● | ◆ | ● | ◆ |  | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ◆ | ★ | ◆ | ○ | ○ | ● | ○ | ● | ● | ★ | ○ | ○ | ● | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ★ |  | ● | ● |  | ● |  | ● |  |  | ○ | ○ | ◆ |  |  | ● | ★ | ○ | ● |  | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○
105 | Disinfectant | T-Spray | Spray | Low | Pharmaceutical Innovations, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○(20) | ○(20) | ○ | ○ | ○ |  | ○ | ○ | ○ |  |  | ○ | ○ |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | X | ○ | ○(20) | ○(20) | ○ | ● | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○(20) | ○(20) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [5 empty] | ○ | ○(20) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(20) | ○(20) |  | ○(20)
105 | Disinfectant | T-SprayⅡ | Spray | Low | Pharmaceutical Innovations, Inc. (USA) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  |  | ○ | ○ | ○(20) | ○(20) | ○ | ○ | ○ |  | ○ | ● | ● |  |  | ● | ○ |  |  | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | X | ● | ○(20) | ○(20) | ● | ● | ● | ● | ○ | ● | ● | ● | ○ | ○ | ● | ● | ● | ● | ● | ● | ● |  | ○(20) | ○(20) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | ○ | ○ | [5 empty] | ○ | ○(20) | ● | ● | ○ | ● | ○ | ● | ○ | ○ | ○ | ○ | ● |  | ○ | ● | ● | ○ | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ | ○ | ○(20) | ○(20) |  | ○(20)
4 | Disinfectant | Tuffie large detergent wipes | Wipes | Low | Vernacare | Ethylan CPG, Betain 810 | X | ○ | ○ | ○
15 | Disinfectant | U Klean | Wipes | Low | Caremate (Korea) | Quaternary Ammonium(N-Alkyl) | [33 empty] | ● |  |  |  |  | ● | ● | [8 empty] | ● | [5 empty] | ● | [35 empty] | ● | ● |  | ● |  | ● | [10 empty] | ● | bar: ◆ ◆ to ◆ ◆; ● ●; ● ●
72 | Disinfectant | UltrOx HLD | Liquid | High | CIVCO (USA) | Hydrogen Peroxide |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ | ○ | ○ | ● | ● |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ● | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | X | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ● | ● | ● |  |  | ● | ○ | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | [15 empty] | ● | ◆ | ◆
66 | Disinfectant | UMONIUM38® NEUTRALIS SPRAY | Spray | Low | Laboratoire Huckert’s International (Belgium) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | ○ | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● |  | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | [35 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ○ ○ to ● ●; ★ ★
67 | Disinfectant | UMONIUM38® Probes | Wipes | Low | Laboratoire Huckert’s International (Belgium) | Quaternary Ammonium(N-Alkyl) | X | X | X | ○ | ● | ● | ● | ● | ● |  |  |  |  | ● | ● | ● |  | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  | ● | ● |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | [34 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ○ ○ to ● ●; ★ ★
67 | Disinfectant | Velox Wipes NA | Wipes | Low | Medisept (Poland) | Quaternary Ammonium(N-Alkyl) | X | X | ○ | X | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | [35 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ○ ○ to ● ●; ○ ○ to ★ ★
52 | Disinfectant | VERT 2 GO ED | Liquid | Low | Wood Wyant (Canada) | Quaternary Ammonium(N-Alkyl) |  |  |  |  | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  |  |  | ○ | ○ |  |  | ○ |  |  |  |  | ○ | ○ | ○ | ○ | ○ | ○ | ○ |  | ○ |  | ○ |  |  | ○ |  |  | ○ | ○ | [34 empty] | ○ | ○ |  | ○ |  | ○ |  |  |  |  | ○ |  |  | ○ | ○ |  | ○ | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ● ● to ● ●; ★ ★
3 | Disinfectant | Virex® II 256 | Liquid | Low | Diversey (Canada) | Quaternary Ammonium(N-Alkyl) |  | X | ○ | ○
62 | Disinfectant | Virkon S* | Powder | High | LANXESS Corporation (Australia) | Potassium peroxymonosulfate | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | ○ | [8 empty] | ○ | [6 empty] | ○ | ○ |  | ● |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ○ | ○ | ○ | ● | ● |  | ○ |  | ● | ● | ● | ○ |  |  | ● | ● | [34 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ○ | ○ |  | ○ | ○ |  | ● | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ○ ○ to ● ●; ★ ★
58 | Sterilization system | V-PRO | Vapor | Sterilization | Steris (USA) | Hydrogen peroxide |  |  |  |  | ★ | ★ | ★ | ★ | ★ | [8 empty] | ★ | [6 empty] | ★ | ★ |  |  |  | ★ | ★ |  |  | ★ |  |  |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  | ★ |  | ★ |  |  | ★ | ★ | ★ | ★ |  | ★ | ★ | ★ |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ |  |  |  | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ★ | ◆ | ◆/★ |  |  |  |  | ★ |  | ★ | ★ |  | ★ |  | ★ |  |  |  |  | ★ |  |  | ★ | ★ |  | ★ |  |  |  |  | ★ |  |  |  | ★
0 | Gel | Vue Ultrasound Gel | Gel | Gel | Optimum Medical (U.K) | Triethanolamine
14 | Disinfectant | V-Wipes | Wipes | Low | Whiteley Corporation (Australia) | Alcohol |  |  |  |  | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ○ | ○ |  |  | ● | ● | [63 empty] | ○ | bar: ◆ ◆ to ◆ ◆; ★ ★
1 | Gel | Wavelength | Gel | Gel | National Therapy Products Inc. (USA) | NA | [92 empty] | ○
103 | Disinfectant | Wavicide®-01 | Liquid | High | Medical Chemical Corp. (USA) | Glutaraldehyde | X | X | ○ | ○ | ● | ● | ● | ● | ● | ○ | X | X |  | ○ | ○ | ○ | ○(50) | ● |  |  | ○(50) | ○(50) | ○(50) | ○(50) | ● | ● |  | ○ |  | ● | ● | ○ |  | ● | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | ● | ● | ○ | ● | ○ | ● | ○(50) | ○(50) | ● |  |  | ● | ○ | ● | ● | ● | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ● |  | ○(50) | ○(50) | ● | ● | ● | ● | ● | ● | ● | ● | ● | ● | x | x | [5 empty] | ● | ○(50) | ● | ● | ○(50) | ● | ○ | ● |  |  |  |  | ● |  | ○ | ● | ● |  | ● | ○ | ○ | ○ |  | ○ | ○ | ○ | ● | ○ | ○ | ○(50) | ○(50) |  | ○(50) | ○ | ● | ● | ◆
70 | Cleaner | Wet Wipe Ethanol 70% Disinfection | Wipes | Cleaner | Wetwipe (Denmark) | Ethanol 70% | X | ○ | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | [35 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● |  |  |  |  | ● | ● | ● | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ○ ○ to ○ ○; ○ ○ to ★ ★
70 | Disinfectant | Wet Wipe Triamin Disinfection | Wipes | Low | Wetwipe (Denmark) | Quat Ammonium | ○ | ○ | X | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | [35 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● |  |  |  |  | ● | ● | ● | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ○ ○ to ○ ○; ○ ○ to ★ ★
67 | Disinfectant | Wip' Anios Excel | Wipes | Low | Anios (France) | Quaternary Ammonium(N-Alkyl) | ○ | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● |  |  | ● | ● |  |  | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● | ● | ● | ● | ● | [33 empty] | ● | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ◆ ◆ to ★ ★; ● ● to ● ●; ○ ○ to ○ ○; ★ ★
60 | Disinfectant | Wip' Anios Spor'Active | Foam | High | Anios (France) | Peracetic Acid | X | ○ | ○ | ○ | ● | ● | ● | ● | ● | [8 empty] | ● | [6 empty] | ● | ● |  |  |  | ● | ● |  |  | ● |  |  |  |  | ● | ● | ● | ● | ● | ● | ● |  | ● |  | ● | ● | ● | ● |  |  | ● | ● | [34 empty] | ● | ● |  | ● |  | ● |  |  |  |  | ● |  |  | ● | ● |  | ● | bar: ◆ ◆ to ◆ ◆; ◆ ◆ to ★ ★; ○ ○ to ● ●; ★ ★
4 | Disinfectant | Yuhanrox regular (Yuhanrox regular 1: water 10) | Liquid | Low | Yuhan clorox (Korea) | Sodium Hypochlorite | ○ | ○ | ○ | ○
28 | Disinfectant | 닥터소독티슈 | Wipes | Low | Carepharm (Korea) | Quaternary Ammonium(N-Alkyl) | [29 empty] | ● | ● | [9 empty] | ● | ● | ● | ● | ● |  | ● |  |  |  |  | ● | [48 empty] | ● |  |  | ● | ● | bar: ◆ ◆ to ★ ★; ○ ○ to ● ●
4 | Disinfectant | 绿先锋消毒剂 | Liquid | High | Techgreen (China) | Chlorine dioxide | bar: ○ ○; ● ●; ○ ○ to ○ ○
↑insert new
System
x = Not compatible(DO NOT USE)
○ = Compatible
Transducer
x = Not compatible(DO NOT USE)
○ = Compatible (From Lens to Housing)
● = Compatible (From Lens upto 50cm away from the connector)
Blank = Untested.(DO NOT USE)
★ = Staining may occur on surface of transducer; however, the acoustic performance and image quality are not affected.
○(5) = Must not be used longer than 5 minutes / ○(10) = Must not be used longer than 10 minutes
○(15) = Must not be used longer than 15 minutes / ○(20) = Must not be used longer than 20 minutes
○(25) = Must not be used longer than 25 minutes / ○(30) = Must not be used longer than 30 minutes
○(50) = Must not be used longer than 50 minutes
All transducers are under tested in IPx 7 Criteria.
◆ Compatible and Validated for disinfection method
◆/★ = Compatible and Validated but staining may occur on the exterior; however, the acoustic performance and image quality are not affected.
* Be sure to check the manufacturer's instructions for the dilution strength and method for using powders including Perasafe.
* V-PRO products include V-PRO1, V-PRO1 Plus, V-PRO Max, V-PRO Max2, V-PRO 60 sterilizer.
* Be sure to follow the manufacturer's instructions before using V-pro sterilization.
* Be sure to check the manufacturer's instructions before using Antigermix disinfection system to prevent discoloration.